SCL-22.2	A	UR.3. Use Requirements, UR.3.1.2 Minimum Load for Cannabis products.

OWM suggested versions:

UR.3.1.2. Minimum Load for Cannabis Products. - Scales used to weigh Cannabis and Cannabis products that have a total weight of 113 grams (4 ounces) or less shall be Class II and the minimum load shall be 100 e.

[Nonretroactive as of January 1, 20XX, to become retroactive as of January 1, 20XX]

(Added 20XX)

or

UR.3.15. Minimum Load for Cannabis Products. - Scales used to weigh Cannabis and Cannabis products that have a total weight of 113 grams (4 ounces) or less shall be Class II and the minimum load shall be 100 e.

[Nonretroactive as of January 1, 20XX, to become retroactive as of January 1, 20XX]

(Added 20XX)

Current version, properly formatted:

UR.3.1.2. Required Minimum Loads for Cannabis Products. - Scales manufactured after January 1, 1986 that are used to weigh Cannabis and Cannabis products that have a total weight of 4 ounces (113 grams) or less shall be Class II and the minimum load shall be 100 e.

[Nonretroactive as of January 1, 20XX, to become retroactive as of January 1, 20XX]

Note: The use of italicized text in the references to “Cannabis” is only to denote its proper taxonomic term; the italicized font does not designate a “nonretroactive” status as is the convention used throughout NIST Handbook 44.

(Added 20XX)

The OWM version does not include:

- The phrase “manufactured after January 1, 1986”, because it has been identified that this could allow the use of unmarked scales to weigh cannabis

- The note, previously part (a), because the item is proposed as a nonretroactive requirement, and the entire paragraph should be italicized. A note may be appropriate if or when the paragraph becomes retroactive.

- The word “required” in the title of the paragraph.  Other paragraphs with similar requirements, such as UR.3.1.1., UR.3.8., and UR.3.9., do not include the word 'required' in the title or body of the paragraph.

An alternative number for the paragraph is also included, as adopting the item with a paragraph number of UR.3.1.2. would add to a conflict that already exists between the recommended minimum loads in UR.3.1. Recommended Minimum Load, and the minimum loads (not recommended but required) in UR.3.1.1. Minimum Load, Grain Dockage Determination.

The phrase “5 years after adoption” is also removed from the designation dates for nonretroactivity and retroactivity in both versions, as this is not the convention used in NIST Handbook 44.  The duration for nonretroactivity should be identified in the justification of the item.  The effective dates for nonretroactivity and retroactivity would be added upon adoption of the item. For example, if adopted in 2026, the designation would specify Nonretroactive as of January 1, 2027, to become retroactive as of January 1, 2032, if the 5-year spread is supported.

The weights and measures community should further contemplate whether the value of 100 e is the appropriate value, as any Cannabis sales of less than 1 g would require an e of 0.005 g or less.  There is no technical reason provided to support the 100 e value.  If a technically sound evaluation is completed, and the weights and measures community determines that 100 e is the correct value, OWM will support this value.